REAGEN™

Pirlimycin ELISA Test Kit

RND99095

### Product Description

REAGEN™ Pirlimycin ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of pirlimycin in milk. Pirlimycin is a lincosamide antibiotic. Under the tradename Pirsue, it is used inthe treatment of mastitis in cattle. Pirlimycin is active against gram-positive bacteria, specifically Staphylococcus aureus and coagulase negative species of Staphylococcus and Streptococcus. It is bacteriostatic and acts by inhibiting bacterial protein synthesis via binding with the 50S subunit of the ribosome.

Pirlimycin ELISA Test Kit enables international and government regulatory agencies, food

manufacturers and processors, as well as quality assurance organizations, to detect pirlimycin in various sample types and to satisfy customer concerns about food safety.

The unique features of the kit are:

·Rapid (10 - 40 minutes), cost-effective extraction methods offering high recovery rates

·High sensitivity (5.0 ng/g or ppb)

·Quick ELISA assay (less than 1.5 hours regardless of number of samples)

·High reproducibility

### Procedure Overview

GENERAL INFORMATION

The method is based on a competitive colorimetric ELISA assay. The Pirlimycin antibody has been coated in the plate wells. During the analysis, sample and Pirlimycin -horseradish peroxidase (Pirlimycin -HRP) conjugate are added to the wells for incubation. If the Pirlimycin residue is present in the sample, it will compete for the Pirlimycin antibody, thereby preventing the Pirlimycin-HRP from binding to the antibody attached to the well. The resulting color intensity, after addition of the HRP substrate (TMB), has an inverse relationship with the Pirlimycin residue concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ The Pirlimycin ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Pirlimycin Antibody-Coated Plate | 1 x 96-well Plate(8 wells x 12 strips) | 2-8C |
| Pirlimycin Standards: Negative control (white cap tube) 5ng/mL (yellow cap tube) 50ng/mL (orange cap tube) 250 ng/mL (pink cap tube) 750ng/mL (purple cap tube) 2250ng/mL (blue cap tube) 5000 ng/mL (spiking, red cap tube) | 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL 0.8 mL | 2-8C |
| Pirlimycin - HRP-Conjugate | 6ml | 2-8C * |
| PLM Milk Extraction Buffer | 14 mL | 2-8C |
| PLM Milk Balance Buffer | 20mL | 2-8C |
| 20X Wash Solution | 28 mL | 2-8C |
| Stop Buffer | 14mL | 2-8C |
| TMB Substrate | 12 mL | 2-8C |

·If you are not planning to use the kit for over 1 months, store Pirlimycin Standard and Pirlimycin -HRP Conjugate at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Milk | 10 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Pirlimycin | 100 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Vortex mixer(e.g.Genie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

### Warnings and Precautions

- The standards contain Pirlimycin . Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using distilled or deionized water since water quality is important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

### Milk

1. Add 1.5 mL of milk sample into a plastic tube. Add 150 uL of PLM Milk Extraction Buffer to each sample.

2. Vortex vigorously for 30 seconds.

3. Centrifuge at 4,000 x g for 10 minutes.

4. Take 660uL of supernatant and add 200 uL of PLM Milk Balance Buffer, mix well by vortexing.

5. Use 100uL of the sample for the assay.

Note: Dilution factor: 1.43

Pirlimycin ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions”

section on page 3. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Buffer concentrate with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Pirlimycin -HRP Conjugate | 50L | 1.2mL |
| 1X Wash Solution | 2.0 mL | 48mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

- Add 100L of each Pirlimycin Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

- Add 100L of each sample in duplicate into different sample wells.

- Add 50L of Pirlimycin -HRP Conjugate to each well, and mix well by gently rocking the plate manually for 1 min.

- Incubate the plate for 30 minutes at room temperature (20 – 25(C / 68 – 77(F).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. ( Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubation for 20 minutes in room temperature, add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Pirlimycin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical Pirlimycin standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us