REAGEN™

Ribavirin ELISA Test Kit

RND99081

### Product Description

REAGEN™ ribavirin ELISA Test Kit enables government agencies, food manufacturers and quality assurance organizations to detect ribavirin in meat,feed to satisfy customer concerns about food safety.

The unique features of the kit are:

- The extraction of ribavirin from samples can be finished in 30 minutes with a high recovery rate of 80 - 115%.

- A quick ELISA assay (less than 3 hours regardless of number of samples).

- High reproducibility.

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ribavirin ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Ribavirin -Coated Plate | 1x96-well Plate (8wells x12strips) | 2-8C |
| Ribavirin Standards: Negative control 0.1ppm 0.3ppm 0.9ppm 2.7ppm 8.1ppm | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | 2-8C |
| HRP-Ab2# Conjugate(red cap tube) | 12mL | 2-8℃ |
| Ribavirin Ab 1#(green cap tube) | 7mL |  |
| 2X Sample Diluent (blue cap tube) | 50 mL |  |
| 20X Wash Solution (white cap tube) | 40 mL |  |
| Stop Buffer (yellow cap tube) | 7 mL |  |
| Substrate A (brown cap tube) | 7 mL |  |
| Substrate B (black cap tube) | 7 mL |  |

### Sensitivity

| Sample Type | Detection Limit (ppm) |
| --- | --- |
| milk | 2 |
| tissue | 0.15 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| ribavirin | 100.0 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Genie Vortex mixer from VWR)

- 10, 20, 100 and 1000 uL pipettes

- Multichannel pipette: 50-300 L (Optional)

### Warnings and Precautions

- The standards contain ribavirin. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. HRP CONJUGATES AND PLATES ARE KIT-AND LOT-SPECIFIC.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25℃ / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X Sample Diluent:

Mix 1 volume of 2X Sample Diluent with 1 volumes of water

### Milk

    - Get 1ml sample and Centrifuge for 10 minutes at 3,000 x g ,remove the fat.

    - Transfer 50 µL of the liquor to 950µL Sample Diluent:

    - Use 50 µL in the assay.

Note: Dilution factor: 20.

### Tissue

    - Homogenize a reasonable amount of sample with a suitable mixer.

    - Weigh out 2 g of the homogenized sample and add 0.5g NaCl,5 mL of acetonitrile.

    - Vortex for 1 minute at maximum speed.

    - Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25℃ / 68 – 77F).

    - Transfer 3mL of the supernatant  and evaporate in mi-Vac Evaporator at 60℃ or dry by blowing nitrogen gas at 60℃.

    - Dissolve in 0.5 mL of 1X Sample Diluent and 0.5ml n-hexane. Vortex for 1 minute to resuspend sample

    - Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25℃ / 68 – 77F).remove the n-hexane.

    - Use 50 µL in the assay.

Note: Dilution factor: 0.5.

Ribavirin ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25℃ / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

1.	Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| HRP Conjugate | 100 uL | 2.4 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 50 uL | 1.2 mL |
| TMB Substrate | 50 uL | 1.2 mL |

    - Add 50 uL of each Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

    - Add 50 uL of each sample in duplicate into different sample wells.

    - Add 50 uL of Ribavirin Ab 1# and mix well by gently rocking the plate manually for 1 minute.

    - Incubate the plate for 30 min at room temperature (20 – 25℃/ 68 – 77F) in the dark.

    - Wash the plate 4 times with 250 uL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100 uL of  HRP,Incubate the plate for 30 min at room temperature (20 – 25℃ / 68 – 77F) in the dark.

    - Wash the plate 4 times with 250 uL of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 50 uL of substrate A.Add 50 uL of  substrate B Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. Incubate for 15 minutes at room temperature (20 – 25C / 68 – 77F) in the dark. (Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

    - After incubation, add 50 uL of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength (Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Ribavirin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve. A special program with Excel functionality,

The following figure is a typical RIBAVIRIN standard curve.

0.1        0.3      0.9        2.7      8.1(PPM)

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us