REAGEN™

Ceftiofur ELISA Test Kit

RND99060

### Product Description

REAGEN™Ceftiofur ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Ceftiofur in milk, urine, tissue (liver, kidney, and muscle), serum, embryonic eggs, and plasma samples.. The features of the kit are:

- High recovery (>80%), low-cost rapid sample prep protocol (10 minutes - 1 hour)

- High sensitivity: 2 ng/gram = 2 ppb (parts per billion)

- Rapid assay protocol (less than 2 hours to process 40 samples in duplicate)

- High reproducibility (interplate and intraplate CV typically <8%)

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest (Ceftiofur) has been coated in the plate wells. The sample to be tested is mixed with a high-affinity Ceftiofur capture protein and added to the well. If the Ceftiofur target is present in the sample, it will compete for the capture protein, thereby preventing the capture protein from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the capture protein bound to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample. The levels of Ceftiofur in the sample are determined by comparison of the color intensity of unknown samples with the standard curve, which is made using the Ceftiofur standard provided in the kit.

### Kit Contents, Storage and Shelf Life

REAGEN™ Ceftiofur ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8°C. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Ceftiofur Coated Plate | 1 x 96-well Plate (8 wells x 12 strips) | 2-8C |
| Ceftiofur powder stock (1000 ng): Empty Vials for Standards: Negative control (white cap tube) 2 ng/mL (yellow cap tube) 8 ng/mL (orange cap tube) 16 ng/mL (pink cap tube) 32 ng/mL (purple cap tube) 64 ng/mL (blue cap tube) | 2 1 1 1 1 1 1 | 2-8C * |
| 10X Ceftiofur Capture Protein | 1.8 mL | 2-8C * |
| Milk Standard Dilution Buffer 1 (Optional) | 15 mL | 2-8C * |
| 200X HRP-Conjugated Antibody #2 | 100 L | 2-8C * |
| Antibody #2 Diluent ** | 20 mL | 2-8C |
| 10X Sample Extraction Buffer** | 2 x 25 mL | 2-8C |
| 20X Wash Solution ** | 28 mL | 2-8C |
| Stop Buffer ** | 14 mL | 2-8C |
| TMB Substrate ** | 12 mL | 2-8C |

* 	If the kit is to be stored for over two weeks before use, store Standard Dilution Buffer1, Ceftiofur powder stock, 10X Ceftiofur Capture Protein and 200X HRP-Conjugated Antibody #2 at -20C in a freezer that does not cycle to higher temperatures (i.e. a non-frost-free freezer). To avoid repeated freeze-thawing of the capture protein it is strongly recommended to store this component in single-use aliquots.

Sensitivity (Detection Limit)

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Milk, Meat, Tissue, Fish and Urine, Embryonic Eggs | 20 |
| Serum and Plasma | 40 |

Specificity (Cross-Reactivity)

| Ceftiofur | Cross-Reactivity (%) |
| --- | --- |
| Ceftiofur | 100 |
| Cefoperazone | 102 |
| Oxacillin | 193 |
| Cloxacillin | 195 |
| Amoxicillin | 208 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Genie Vortex mixer from VWR)

- 10, 20, 100 and 1000 (L pipettes

- Multi-channel pipette: 50-300 (L (Optional)

- Distilled deionized water

- Disposable plastic reservoirs

- Paper towels

### Warnings and Precautions

- The standards contain Ceftiofur. Handle with care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT-AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20 – 25C / 68 – 77F Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 10X Sample Buffer with 9 volumes of distilled water.

- Preparation of 70% Methanol

Mix 7 volumes of methanol with 3 volumes of 1X Sample Extraction Buffer.

- Preparation of 35% Methanol

Mix 7 volumes of methanol with 13 volumes of 1X Sample Extraction Buffer.

Meat/Fish Tissue/Embryonic Eggs

    - Homogenize a reasonable amount of sample with a suitable mixer until the sample has a consistency like paste. Store sample at -20°C until analyzed (i.e. in a non-frost-free freezer).

    - Weigh out 1 g of the homogenized sample and mix with 4 mL of 70% methanol, vortex for 3 minutes or shake for 20 minutes with a shaker.

    - Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

    - Dilute 0.5 mL of the obtained supernatant with 0.5 mL of 1X Sample Extraction Buffer.

    - Use 75µL of the diluted supernatant per well in the test.

Note: Dilution Factor: 10.

### Milk

For fat-free milk, dilute the milk sample 1:9 with 35% methanol (e.g. 20 L of milk + 180 L of 35% methanol). Use 75 L of the diluted sample per well for the assay.

For regular milk with fat, centrifuge the milk sample at 4,000 rpm for 10 minutes and discard the upper fat layer. A toothpick, pipette tip, or weighing spatula is generally adequate to drag the fat layer to the side or pierce an opening through which you may obtain the separated milk. Dilute the milk sample 1:9 with 35% methanol (e.g. 20 L of milk + 180 L of 35% methanol). Use 75 L of the diluted sample per well for the assay.

Note: Dilution factor:10.

### Serum/Plasma

    - Centrifuge the serum/plasma sample at 4,000 rpm for 15 minutes.

    - Dilute the supernatant 1:19 with 35% methanol (e.g. 10 L of serum/plasma + 190 L of 35% methanol).

    - Use 75 L of the diluted sample per well for the assay.

Note: Dilution factor: 20.

### Urine

    - Centrifuge the urine sample at 4,000 rpm for 10 minutes.

    - Dilute the supernatant 1:9 with 35% methanol (e.g. 20 L of urine + 180 L of 35% methanol).

Note: Dilution factor: 10.

CEFTIOFUR ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3.Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of Ceftiofur Standards

Add 1 mL of 35% methanol (see page 4 for preparation) to the stock vial (1000 ng) and mix by vortexing for ~30 sec to obtain 1000 ppb stock.  Centrifuge the tube briefly to re-collect the liquid. To make the working standards, serially dilute the 1000 ppb standard stock vial with 35% methanol in the provided empty standard vials as described in the following table. Make sure to mix each standard vial by vortexing thoroughly and briefly centrifuge before removing aliquots.

*For Preparing Standards for use with milk ONLY: Use the Milk Standard Dilution Buffer 1, instead of 35% methanol, to make the standard dilutions after the preparation of 1000 ppb stock.

| Standards | Ceftiofur Source | Volume of Ceftiofur Solution to add from Source Tube | Volume of Diluent: (35% methanol or Milk Standard Dilution Buffer 1*) |
| --- | --- | --- | --- |
| 64 ppb vial | 1000 ppb stock vial | 64 L | 936 L |
| 32 ppb vial | 64 ppb vial | 500 L | 500 L |
| 16 ppb vial | 32 ppb vial | 500 L | 500 L |
| 8 ppb vial | 16 ppb vial | 500 L | 500 L |
| 2 ppb vial | 8 ppb vial | 200 L | 600 L |
| Negative Control vial | N/A | 0 L | 500 L |

Note: The 1000 ppb stock vial can be divided into single-use aliquots and stored at -20C in a non-frost-free freezer for 1 month. The working standards must be freshly prepared before use in the ELISA. After using the standards, empty the standard vials and store them at 4C. The empty vials can be re-used for preparing standards for subsequent assays.

- Preparation of 1X Ceftiofur Capture Protein

Mix 1 volume of 10X Ceftiofur Capture Protein with 9 volumes of 1X Sample Extraction Buffer.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 200X HRP-Conjugated Antibody #2 with 199 volumes of Antibody #2 Diluent.

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| 1X Ceftiofur Capture Protein | 100 L | 2.4 mL |
| 1X HRP-Conjugated Antibody #2 | 150 L | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

    - Add 75 L of each Ceftiofur Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

    - Add 75 L of each sample in duplicate into different sample wells.

    - Add 100 L of 1X Ceftiofur Capture Protein and mix well by gently rocking the plate manually for 1 minute.

    - Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F). 

    - Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 150 L of 1X Antibody #2 solution. Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F) (  Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

    - Wash the plate 3 times with 250 of L 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (  Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

    - After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength (( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings)

### Ceftiofur Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical Ceftiofur standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us