REAGEN™

Olaquindox ELISA Test Kit

RND99051

### Product Description

REAGEN™ Olaquindox ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Olaquindox in meat/fish/shrimp/egg and feed.

The unique features of the kit are:

- High recovery (＞80%) .

- Sensitivity:(0.2ppb)

- A quick ELISA assay.

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The Olaquindox has been coated in the plate wells. During the analysis, sample is added along with the Olaquindox antibody. If the Olaquindox residue is present in the sample, it will compete for the Olaquindox antibody, thereby preventing the antibody attached to the well. The resulting color intensity, after addition of the HRP substrate (TMB), has an inverse relationship with the Olaquindox residue concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Olaquindox ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C*. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Olaquindox -Coated Plate | 96-well plate (8 wells x 12 strips) | 2-8C |
| Olaquindox Standards: Negative control 0.2ng/mL 0.4ng/mL 0.8 ng/mL 1.6ng/mL 3.2ng/mL | 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL 1.8 mL | 2-8C |
| 100X Olaquindox conjugate | 100uL | 2-8C |
| 10X Extraction & Wash Solution | 50 mL |  |
| 10x Conc. Buffer | 50 mL |  |
| Stop Buffer | 13 mL |  |
| TMB Substrate | 12 mL |  |

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| meat/fish/shrimp/egg | 1 |
| Feed | 40 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Olaquindox | 100 |
| Quinolone | 147 |
| Olaquindox metabolite | <0.5 |
| 2- Quinoxaline carboxylic acid | <0.5 |
| Mequindox | 69 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni Tissue Master Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- 0.01M Hydrochloric acid solution

### Warnings and Precautions

- The standards contain Olaquindox .Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. Olaquindox-HRP Conjugates and plates are kit and lot-specific.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X Extraction & Wash Solution:

Mix 1 volume of 10X Extraction & Wash Solution with 9 volumes of distilled water.

Preparation of 1X Conc. Buffer:

Mix 1 volume of 10X Conc. Buffer with 9 volumes of distilled water.

Preparation of 1X Olaquindox conjugate:

Mix 1 volume of 100X Olaquindox conjugate with 99 volumes of 1X Conc. Buffe.

- Preparation of 0.01M Hydrochloric acid solution:

- Carefully remove 83ul 36% of concentrated hydrochloric acid solution constant volume in 100 ml of deionized water.According to the need can be enlarged or reduced in proportion.

### Meat/fish/shrimp/egg

- Weigh 2g of the ground sample, add 3ml of 1X Conc. Buffer, vortex vigorous for 30 s .

- The sample can be Incubate for 15 minutes in a 70C water bath（Avoid light）.

- Centrifuge the sample at 4,000 x g for 15 minutes.

- Transfer 200 L of the supernatant,add 200 L of 1X Extraction & Wash Solution, vortex vigorous for 1 minute.

- Use 50 L for the assay.

Note: Dilution factor: 5.

### Feed

1.  Weigh 1g of the ground feed sample, add 10 ml of 0.01M Hydrochloric acid solution, vortex vigorous for 10 minute .

2.   Centrifuge at 4,000 x g for 15minutes at room temperature.

- Transfer 50 L of the supernatant into a new vial and add 0.95mL of 1X Conc.Buffer, vortex vigorously for 1 minute..

4.   Use 50 L for the assay.

Note: Dilution factor: 200

OLAQUINDOX ELISA TEST KIT PROTOCOL

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| 1X Olaquindox conjugate | 50 L | 1.2 mL |
| 1X Wash Solution | 1 mL | 24 mL |
| Stop Buffer | 100L | 2.4 mL |
| TMB Substrate | 100L | 2.4 mL |

- Add 50 L of each Olaquindox standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

- Add 50 L of each sample in duplicate into different sample wells.

- Add 50 L of 1X Olaquindox conjugate and mix well by gently rocking the plate manually for 1 min.

- Incubate the plate for 30 minutes at 22.5±2.5°C .

Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

REAGEN LLC

Add.: 102CommerceDr.Unit8, Moorestown,NJ08057

Tel.:856-727-0250         FAX.: 856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us

- Add 100 L TMB .Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 20 minutes at 22.5±2.5°C, add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### OLAQUINDOX Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.