REAGEN™

Avermectins/ Ivermectin ELISA Test Kit

RND99032

### Product Description

REAGEN™ Avermectins/ Ivermectin ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Abamectin in meat, animal tissue, Vegetables and Cereal.

The unique features of the kit are:

- The extraction of Avermectins from milk can be finished with a high recovery rate of 75 - 105%.

- High reproducibility.

- A quick ELISA assay (less than 2 hours regardless of number of samples).

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Avermectins/ Ivermectin ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Avermectin-coated Microtiter Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8C |
| Ivermectin Standards: Negative control 0.5 ng/mL 1.0 ng/mL 2 ng/mL 4 ng/mL 8 ng/mL 1200 ng/mL (spiking, optional) | 1.0mL 1.0mL 1.0mL 1.0mL 1.0mL 1.0mL 1.0mL | 2-8C * |
| Avermectin Antibody #1 | 12ml | 2-8C |
| 100X HRP-Conjugated Antibody #2 | 250 L |  |
| Antibody #2 Diluent | 20 mL |  |
| 20X Wash Solution | 28mL |  |
| Stop Buffer | 14 mL |  |
| TMB Substrate | 12 mL |  |
| Avermectin Clean Up Mix (Optional) Milk Extraction Buffer I (Optional) Milk Extraction Buffer II(Optional) Avermectin Balance Buffer I (Optional) Avermectin Balance Buffer II (Optional) Liver Extraction Buffer (Optional) 10X Sample Extraction Buffer | 25g 10 mL 28 mL 10 mL 62g 20 mL 25mL |  |

* 	If you are not planning to use the kit for over 1 month, store Ivermectin Standard , Avermectin Antibody #1, 100X HRP-Conjugated Antibody #2 at -20C preferably or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Meat/Fish/Shrimp | 2.5 |
| Liver | 10 |
| Milk | 2.5 |
| Urine/Serum | 2.5 |
| Water | 2.5 |

### Specificity

| Molecules | Cross-Reactivity (%) |
| --- | --- |
| Ivermectin | 100% |
| Emamectin | 100% |
| Abamectin | 113% |
| Eprinomectin | 100% |
| Doramectin | <1% |
| Moxidectin | <1% |
| Milbamycin | <1% |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Acetonitrile ，ethanol

### Warnings and Precautions

- The standards contain Ivermectin. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT-AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

- Preparation of 35% Ethanol/ Sample Extraction Buffer

Mix 6.5 volumes of 1X Sample Extraction Buffer with 3.5 volumes of 100% ethanol

- Preparation of 1X Avermectin Balance Buffer

Add 10 mL of the Avermectin Balance Buffer I and 62 g of Avermectin Balance Buffer II (Powder) to a 250-mL plastic bottle. Add 190 mL of distilled water to the bottle, mix well to obtain a homogeneous solution.

### Meat/Fish/Shrimp

    - Homogenize a reasonable amount of sample with a suitable mixer.

    - Weigh out 2 g of the homogenized sample. Add 6 mL of acetonitrile, 2 mL of n-hexane, and 1 g of anhydrous MgSO4, vortex 1 min at maximum speed and then at rotorack for 15 minutes.

    - Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

    - Remove the hexane layer as completely as possible. Take 1.0 mL of the supernatant into a new 2 mL tube containing 250 mg of Avermectin Clean Up mix.

    - Vortex vigorously for 1 minute and let it stand at room temperature for 5 minutes.  Centrifuge the sample for 5 minutes at maximum speed at room temperature (20 – 25C / 68 – 77F).

    - Take 400 L of supernatant to a 2 mL tube and use an evaporator to dry the sample at 50C under reduced pressure.  Alternatively, the sample can be dried by blowing nitrogen gas in a 50C waterbath.

    - Add 500 L of 35% Ethanol/Sample Extraction Buffer to the dried tube and vortex vigorously for 2 minutes.

    - Use 50 uL of the suspension per well in the ELISA test.

Note: Dilution factor: 5.

### Liver

    - Weigh 2 g homogenized liver sample, add 6 mL of acetonitrile and 2.5 mL of 1X Avermectin Balance Buffer

    - Vortex for 3 minutes at maximum speed or shake for 20 minutes at a multi-tube vortexer or shaker.

    - Add 200 mL of Liver Extraction Buffer, vortex for 1 minute.

    - Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F), transfer 200 mL of the supernatant to a new tube. Use a rotary evaporator to dry the sample in a 50-60°C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50-60°C water bath.

    - Add 1mL of 35% Ethanol/Sample Extraction Buffer to the dried tube and vortex vigorously for 1 minute.

    - Use 50 mL of the sample for the assay.

Note: Dilution factor: 20.

### Milk

- For the regular milk with fat, centrifuge the milk sample at 10,000 x g for 5 minutes or 4,000 x g for 30 minutes, discard the upper fat layer. For fat-free milk, skip this step.

- After remove the fat, take 2 mL of milk sample, add 6 mL of acetonitrile and 0.1 mL of Milk Extraction Buffer I, shake 5 seconds after Milk Extraction Buffer I is added to all the samples, and then add 0.35 mL of Milk Extraction Buffer II, manually vortex at maximum speed for 2 minutes or shake for 10 minutes in a multi-vortexer or shaker.

- Centrifuge at 4,000 x g for 10 minutes.

- Transfer 0.4 mL of the acetonitrile supernatant (corresponding to 0.1 mL of the original milk sample) into a new vial and use a rotary evaporator to dry the sample in a 50C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50C water bath.

- Dissolve the dried residue in 0.5 mL of 35% methanol/Sample Extraction Buffer by vortexing at maximum speed for 1 minute.

- Use 50 L of the sample per well for the assay.

Note: Dilution factor: 5. The dilution factor can be changed to 1 if transfer 2 mL of acetonitrile supernatant in Step 4 instead of 0.4 mL.

### Vegetable/Cereal

- Homogenize a reasonable amount of sample with a suitable mixer..

- Weigh out 2 g of the homogenized sample. Add 6 mL of acetonitrile, and 1 g of anhydrous MgSO4, vortex 1 min at maximum speed and then at rotorack for 15 minutes.

- Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25°C / 68 – 77°F).

- Take 1.0 mL of the supernatant into a new 2 mL tube containing 250 mg of Avermectin Clean Up mix.

- Vortex vigorously for 1 minute and let it stand at room temperature for 5 minutes.  Centrifuge the sample for 5 minutes at maximum speed at room temperature (20 – 25°C / 68 – 77°F).

- Take 400 mL of supernatant to a 2 mL tube and use an evaporator to dry the sample at 50°C under reduced pressure.  Alternatively, the sample can be dried by blowing nitrogen gas in a 50°C waterbath

- Add 500 uL of 35% Ethanol/Sample Extraction Buffer to the dried tube and vortex vigorously for 2 minutes.

- Use 50 uL of the suspension per well in the ELISA test

Note: Dilution factor: 5.

### Urine/Serum

- Centrifuge the urine sample at maximum speed for 5 minutes.

- Transfer 250 L of supernatant to a new tube, add 1 mL of 35% Ethanol/Sample Extraction Buffer and 25 mg of Avermectin Clean Up mix, vortex for 1 minute, stand at room temperature for 5 minutes.

- Centrifuge at maximum speed for 10 minutes.

- Use 50 L of the supernatant per well for the assay.

Note: Dilution factor: 5.

### Water

- Centrifuge 1 mL of water sample at 4000 x g for 5 minutes.

- Transfer 250 L of supernatant to a new tube, add 1 mL of 35% Ethanol/Sample Extraction Buffer. Mix well.

- Use 50 L of the sample per well for the assay.

Note: Dilution factor: 5.

ABAMECTIN ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2 Diluent.

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Avermectin Antibody #1 | 100 L | 2.4 mL |
| 1X-HRP Conjugate Antibody #2 | 150 L | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

- Add 50 L of each Ivermectin Standard in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

- Add 50 L of each sample in duplicate into different sample wells.

- Add 100 L of Avermectin Antibody #1 to each well and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 150 L of 1X HRP-Conjugated Antibody #2. Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F) (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels (Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Ivermectin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical Ivermectin standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250     FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us