REAGEN™

Chlorpromazine ELISA Test Kit

RND99022

### Product Description

REAGEN™ Chlorpromazine ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of Chlorpromazine in milk,feed,urine,serum and meat,liver and kidney.The unique features of the kit are:

- High recovery (75 - 105%),rapid(10-40minutes), cost-effective extraction methods.

- High sensitivity (0.02 ng/g or ppb) and low detection limit (0.1 ng/g or ppb for meat)

- High reproducibility.

- A quick ELISA assay (less than 2 hours regardless of number of samples).

### Procedure Overview

REAGEN™ Chlorpromazine ELISA Test Kit is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The second antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Chlorpromazine ELISA Test Kit has the capacity acity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C . The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Chlorpromazine-coated Microtiter Plate | 1x 96-well Plate (8 wells x 12 strips) | 2-8C |
| Chlorpromazine Standards: Negative control (white cap tube) 0.02ng/mL (yellow cap tube) 0.04ng/mL (orange cap tube) 0.1 ng/mL (pink cap tube) 0.25ng/mL (purple cap tube) 0.5 ng/mL (blue cap tube) 10 ng/mL (spiking, red cap tube) | 1.0mL 1.0mL 1.0mL 1.0mL 1.0mL 1.0mL 1.0mL | 2-8C 2-8C |
| Chlorpromazine Antibody #1 | 10 mL |  |
| 100XHRP Conjugated Antibody #2 | 250L |  |
| Antibody #2 Diluent | 20 mL |  |
| 20X Wash Solution | 30 mL |  |
| TMB Substrate | 10 mL |  |
| Stop Buffer | 15 mL |  |
| Clean Up Mix | 4g |  |
| 10X PBS | 25 mL |  |

* 	If you are not planning to use the kit for over 1 months, store Chlorpromazine  Antibody #1 and 100XHRP-Conjugated Antibody #2 at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Feed | 2.0 |
| Meat/Chicken/Liver,milk | 0.1 |
| Urine/Serum | 0.2 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Chlorpromazine | 100 |
| Promazine | 105 |
| Triflupromazine | 89 |
| Propionylpromazine | 91 |
| Imipramine | 93 |
| Acepromazine | 86 |
| Amitriptyline | 75 |
| Doxeoin | 74 |
| Trimipramine | 77 |
| Promethazine | 67 |
| Protriptyline | 28 |
| Diphenhydramine | 2 |
| Maprotiline | <1 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Rotary evaporator or nitrogen gas

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

### Warnings and Precautions

- The standards contain Chlorpromazine . Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

1.  Preparation of 1X PBS

Mix 1 volume of the 10X PBS with 9 volumes of distilled water.

2.  Preparation of 35% (or 70%)Methanol/1X PBS

Mix 3.5 (or 7) volumes of methanol with 6.5 (or 3) volumes of 1X PBS.

### Feed

- Homogenize a reasonable amount of feed sample by a suitable mixer.

- Add 20 mL of 0.1 M HCl to 2 g of the homogenized sample. Vortex for 15 minutes at a multi-tube vortexer or shaker.

- Centrifuge for 10 minutes at 2,000 x g at room temperature 20°–25°C (68°–77°F), transfer 1 mL of the supernatant to a new tube, add ~55 L of 1 M NaOH to adjust the pH to 6-8 (the amount of 1 M NaOH could be adjusted based on different feed samples).

- Centrifuge the sample for 10 minutes at 2,000 x g, transfer 0.5 mL of the supernatant to a new tube, add 0.5 mL of 70%Methanol/1X PBS, mix well.

- Use 50L of the sample for the assay.

Note: Dilution factor: 20.

### Meat/Chicken/Liver

- Weigh 2 g homogenized sample and add 6 mL of methanol and 2 mL of hexane. Vortex for 3 minutes at maximum speed or shake for 20 minutes at a multi-tube vortexer or shaker.

- Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25°C/ 68 – 77°F), remove the hexane layer completely and then transfer 1 mL of the supernatant to a new tube.

- Add 50 mg Clean Up Mix and vortex for 30 seconds.

- Transfer 500 μL supernatant to a new tube. Use a rotary evaporator to dry the sample in a 50-60°C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50-60°C water bath.

- Add 250 μL of 35% Methanol/1X PBS, vortex vigorously for 1 minute at maximum speed.

- Use 50 μL of the sample for the assay.

Note: Dilution factor: 2.0.

- Urine/Serum

- Centrifuge 0.5 mL of the urine or serum sample at 4,000 x g for 5 minutes.

- Take out 0.1 mL of supernatant, add 0.4 mL of 35% Methanol/1XPBS, mix well.

- Use 50 μL of the supernatant per well for the assay.

Note: Dilution factor:5.0

- Milk

- For fat-free milk, dilute 200 μL of milk sample with 450 μL of 1XPBS and 350 μL of  methanol. Mix well. Take 50 μL of the diluted sample per well for the assay.

- For regular milk with fat, centrifuge the milk sample at 4,000 rpm for 5 minutes, discard the upper fat layer. Dilute 200 μL of the defatted milk sample with 450 μL of 1XPBS and 350 μL of  methanol. Mix well. Take 50 μL of the diluted sample per well for the assay.

Note: Dilution factor: 5.

CHLORPROMAZINE ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test. All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2 Diluent.

- Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Buffer concentrate with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Chlorpromazine Antibody #1 | 100L | 2.4mL |
| 1X HRP-Conjugated Antibody #2 | 150 L | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

    - Add 50Lof each Chlorpromazine Standards in duplicate into different wells (Add standards to plate only in the order from low concentration to high concentration).

    - Add 50L of each sample in duplicate into different sample wells.

    - Add 100L Antibody #1 to each well and mmix well by gently rocking the plate manually for 1 minute.

    - Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F).

    - Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 150 L of 1X HRP-Conjugated Antibody #2 to each well ,Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F).

    - Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

    - Add 100L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. ( Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

    - After incubation for 15 minutes in room temperature, add 100 L of Stop Buffer to stop the enzyme reaction.

    - Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Chlorpromazine Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical Chlorpromazine standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com   reagenllc@gmail.com

Web.:www.reagen.us