REAGEN™

Nitroimidazole ELISA Test Kit

RND99005

### Product Description

REAGEN™ Nitroimidazole ELISA Test Kit provides a competitive enzyme immunoassay for the quantitative analysis of Nitroimidazole in milk、 milk powder .

The unique features of the kit are:

- High recovery (70% - 95%), and cost-effective extraction methods for various samples.

- A quick ELISA assay (less than 1.5 hours regardless of number of samples).

- High reproducibility.

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The Nitroimidazole antibody has been coated in the plate wells. During the analysis, sample is added along with the Nitroimidazole -horseradish peroxidase (HRP) conjugate. If the target is present in the sample, it will compete for the Nitroimidazole antibody, thereby preventing the Nitroimidazole- HRP from binding to the antibody attached to the well. The resulting color intensity, after addition of substrate((TMB), has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™Nitroimidazole ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C*. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Nitroimidazole AB-coated Microtiter Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8C |
| Non-coated Transfer Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8C |
| 10X Metronidazole Standards: Negative control(white cap tube) 0.2ng/mL(yellow cap tube) 0.6ng/mL(orange cap tube) 2ng/mL (pink cap tube) 6 ng/mL (purple cap tube) 20 ng/mL (blue cap tube) Spiking 10 ng/mL (red cap tube) | 0.3mL 0.3 mL 0.3mL 0.3mL 0.3 mL 0.3 mL 0.8 mL | 2-8C 2-8C |
| 100x HRP-Conjugated | 1bottle | 2-8C * |
| HRP Conjugate Diluent | 15 mL | 2-8C |
| TMB Substrate | 12mL | 2-8C |
| 20X Wash Solution ** | 28 mL | 2-8C |
| Stop Buffer ** | 14 mL | 2-8C |
| Milk Extraction Buffer A | 5 mL | 2-8C |
| Milk Extraction Buffer B | 5 mL | 2-8C |
| Milk Dilution Buffer C | 5 mL | 2-8C |
| Standard Dilution Buffer | 25 mL | 2-8C |

* 	If you are not planning to use the kit for over 1 months, store 10X Standards &  100X Nitroimidazole-HRP in a -20 °C freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Milk | 0.1 |
| Milk Powder | 0.2 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Dimetridazole | ＞100 |
| Metronidazole | 100 |
| Hydroxydimetridazole | 51.3 |
| Hydroxymetronidazole | ＞100 |
| Ronidazole | ＞100 |
| Ipronidazole | ＞100 |
| Hydroxyipronidazole | 72.1 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Tissue Mixer

- Incubator

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

### Warnings and Precautions

- The standards contain Dimetridazole . Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. HRP AND PLATES ARE KIT- AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may ause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

### Milk

1.	Transfer 1.0 mL of milk to a 2-mL micro-centrifuge tube, add 50 μL of Milk Extraction Buffer A and 50 μL of Milk Extraction Buffer B.

2.	Vortex sample for 3 minutes.

3.	Incubate sample for 20 minutes in a 60 °C incubator or heat brick.

4.	Vortex sample for 1 minute.

5.	Incubate sample for 10 minutes in a -17 to -20 °C freezer.

6.	Vortex sample for 30 seconds.

7.	Centrifuge sample for 15minutes at maximum speed around 4,000 x g.

8.	Transfer 440 μL of clear supernatant to a new 2-mL tube containing 60 μL of Milk Dilution Buffer C.

9.	Vortex sample for 2 minutes. Incubate sample at room temperature for 15 mintues. Vortex for an additional minute.

10.  Centrifuge sample for 10 minutes at maximum speed around 4,000 x g.

11.  Use 150μL of the clear supernatant per well for the assay.

Note: Dilution factor = 1.25

·Milk Powder

1.   Weigh 1 g of milk powder into a 15-mL tube. Add 9 mL of distilled water. Vortex for 3 minutes to obtain a homogeneous solution.

2.   Continue with sample preparation according to Step 1 of Milk Protocol above.

Note: Dilution factor = 11.25

Nitroimidazole ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents（store antibody at 4℃ before use） should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Solution with 19 volumes of distilled water.

Preparation of 100X Nitroimidazole -HRP Conjugate

Add 300L HRP Conjugate Diluent to bottle of powder Nitroimidazole-HRP Conjugate invert the bottle up and down for 10 times,leave the tube at room temperature for at least 15 min,invert the tube up and down for another 10 time.Vortex gently 30 seconds.The HRP solution is now 100X.Stored the unused 100X D Nitroimidazole -HRP Conjugate in -20C or in a freezer.

Preparation of 1X Nitroimidazole -HRP Conjugate

Mix 1 volume of 100X Nitroimidazole -HRP Conjugate with 99 volumes of HRP Conjugate Diluent.

1X Working Metronidazole Standards

For each concentration of standard (including the negative control), in a new 2-mL tube, combine 50 L of the 10X Standard with 450 L of the Standard Dilution Buffer. Vortex each standard very well for at least 1 minute. It is recommended to prepare 1X Standards each time an ELISA is performed, however, the 1X Standards are stable for up to 3 days when stored at 4 °C. Final standard concentrations will be: 0, 0.02, 0.06, 0.2, 0.6, & 2 ng/mL.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| 1X Nitroimidazole -HRP Conjugate | 50L | 1.2mL |
| 1X Wash Solution | 2.0mL | 48mL |
| Stop Buffer | 100L | 2.4mL |
| TMB Substrate | 100L | 2.4mL |

1.  Add 150 L of each 1X Working Metronidazole Standards in duplicate into different wells of the Transfer Plate. Add standards to plate only in the order from low concentration to high concentration

2.  Add 150 L of each sample in duplicate into different sample well of the Transfer Plate.

3.  Add 50L of 1X Nitroimidazole-HRP Conjugate to the standards and samples in the Transfer Plate.Using a multi-channel pipette, pipette the content of each strip up-and-down at least 20 times.

4.  Incubate the Transfer Plate for 10 minutes at room temperature (20 – 25 °C / 68 – 77 °F).

5.  Using a multi-channel pipette, transfer 150 L from each well in each strip into the Antibody Plate. Mix the Antibody Plate by tapping it against a hard force-point for 1 minute.

6.  Incubate the Antibody Plate for 2 hours at room temperature (20 – 25 °C / 68 – 77 °F).

7.  Wash the plate 3 times with 250 L 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels. Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps.

8.  Add 100 L of TMB substrate. Incubate the plate for 20 minutes at room temperature (20 – 25°C / 68 – 77°F). Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating. Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the plate while incubating is recommended).

9.   After incubation, add 100 L of Stop Buffer to stop the enzyme reaction.

10.   Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm primary filter and 630 nm differential filter wavelengths.Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Nitroimidazole Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical Nitroimidazole standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com   reagenllc@gmail.com

Web.:www.reagen.us