REAGEN™

Ractopamine ELISA Test Kit

RND99014

### Description

REAGEN™ Ractopamine ELISA Test Kit is a competitive enzyme immunoassay for this beta-agonist drug in meat, animal tissue, milk, feed, urine, serum and plasma. REAGEN™ Ractopamine ELISA Test Kit provides a unique extraction method, which is accurate, rapid, cost-effective and easy to use.

The features of the kit are:

- The extraction of ractopamine from feed, meat, milk, tissue and urine can be finished in 10 - 60 minutes with a high recovery rate of 80 - 120%.

- High reproducibility.

- A quick ELISA assay (less than 2 hours regardless of number of samples).

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Ractopamine-coated Microtiter Plate | 96-well plate(8 wells x 12 strips) | 2-8C |
| Ractopamine Standards: Negative control (white cap tube) 0.04 ng/mL (yellow cap tube) 0.1ng/mL (orange cap tube) 0.25 ng/mL (pink cap tube) 0.6ng/mL (purple cap tube) 1.5 ng/mL (blue cap tube) 100ng/mL(spiking, red cap tube) | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | 2-8C |
| Ractopamine Antibody #1 | 15 mL | 2-8C |
| 100X HRP-Conjugated Antibody #2 | 250 L |  |
| Antibody #2 Diluent ** | 20 mL |  |
| 10X Sample Extraction Buffer ** | 25 mL |  |
| 20X Wash Solution ** | 30 mL |  |
| Stop Buffer ** | 15 mL |  |
| TMB Substrate ** | 12 mL |  |
| Ractopamine Clean Up Mix (optional) | 4 g |  |
| 10X Feed Sample Buffer I (optional) | 20 mL X 2 |  |
| 1X Feed Sample Buffer II (optional) | 20 mL |  |
| 1XFeed Sample Buffer III (optional) | 20 mL |  |

* 	If you are not planning to use the kit for over 1 month, store Ractopamine Antibody #1 and 100X HRP-Conjugated Antibody #2 at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Feed | 0.5 |
| Meat/Liver/Kidney | 0.1 |
| Milk | 0.4 |
| Urine | 0.4 |

### Specificity

| Analytes | Sensitivity (ppb) | Cross-Reactivity (%) |
| --- | --- | --- |
| Ractopamine | 0.025 | 100 |
| Ractopamine Glucuronide A | 0.06 | 51 |
| Ractopamine Glucuronide B | 0.1 | 48 |
| Ractopamine Glucuronide C | 0.03 | 95 |
| Clenbuterol | >10,000 | <0.01 |
| Colterol | >10,000 | <0.01 |
| Mabuterol | >10,000 | <0.01 |
| Isoxsuprine | 5,000 | <0.01 |
| Salmeterol | 5,000 | <0.01 |
| Fenoterol | >10,000 | <0.01 |
| Bitolterol | >10,000 | <0.01 |
| Carbuterol | >10,000 | <0.01 |
| Pirbuterol | >10,000 | <0.01 |
| Terbuterol | >10,000 | <0.01 |
| Timolol | >10,000 | <0.01 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Rotary evaporator or nitrogen gas

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Methanol

- n-Hexane

### Warnings and Precautions

- The standards contain ractopamine. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT- AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temperature (20 – 25C / 68 – 77F) or in a refrigerator before use.

  - Preparation of 1X Feed Sample Buffer I:

Mix 1 volume of 10X Feed Sample Buffer I with 9 volumes of distilled water.

  - Preparation of 1X Sample Extraction Buffer

Mix 1 volume of 10X Sample Extraction Buffer with 9 volumes of distilled water.

### Feed

- Grind and mix a representative sample (according to accepted sampling techniques).

- Weigh out 1 g of the ground sample and place into a suitable container.

- Add 5 mL of 1 X Feed Sample Buffer I to all the samples. Shake the samples for 5 seconds. Then add 250 L of Feed Sample Buffer II to each sample, vortex at maximum speed for 3 minutes manually or vortex in a multi-vortexer for 10 minutes.

- Centrifuge at 4,000 x g for 10 minutes.

- Transfer 250 L of clear supernatant to another tube, dilute with 250 L of 1 X Feed Sample Buffer III.

- Vortex for 15 seconds. Centrifuge at 4,000 x g for 2 minutes.

- Use 50 L of the sample per well for the assay.

Note: Dilution factor: 11.

### Meat/Liver/Kidney

- Weigh 2 g homogenized sample and add 6 mL of methanol and 2 mL of hexane. Vortex for 3 minutes at maximum speed or shake for 20 minutes at a multi-tube vortexer or shaker.

- Centrifuge for 10 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F), remove the hexane layer completely and then transfer 1 mL of the supernatant to a new tube.

- Add 50 mg Ractopamine Clean Up Mix and vortex for 30 seconds.

- Transfer 500 μL supernatant to a new tube. Use a rotary evaporator to dry the sample in a 50-60C water bath under reduced pressure. Alternatively, the sample can be dried by blowing nitrogen gas in a 50-60C water bath.

- Add 250 μL of 1X Sample Extraction Buffer, vortex vigorously for 1 minute at maximum speed.

- Use 50 μL of the sample for the assay.

Note: Dilution factor: 2.

### Milk

- For fat-free milk, dilute the milk sample with 1X Sample Extraction Buffer (1:9) (e.g. 20 μL of milk + 180 μL of buffer). Take 50 μL of the diluted sample per well for the assay.

- For regular milk with fat, centrifuge the milk sample at 4,000 x g for 5 minutes, discard the upper fat layer. Dilute the sample with 1X Sample Extraction Buffer (1:9) (e. g. 20 μL of milk +180 μL of buffer). Take 50 μL of the diluted sample per well for the assay.

Note: Dilution factor: 10.

### Urine/Serum/Plasma

- Centrifuge 0.5 mL of the sample at 4,000 x g for 5 minutes.

- Dilute the supernatant with 1X Sample Extraction Buffer (1:9) (e. g. 20 μL of supernatant + 180 μL of buffer). Take 50 μL of the diluted sample per well for the assay.

Note: Dilution factor: 10.

The urine sample can also be directly used in the ELISA without dilution.

RACTOPAMINE ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 4. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2 Diluent.

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Ractopamine Antibody #1 | 100 L | 2.4 mL |
| 1X HRP-Conjugated Antibody #2 | 150 L | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 L | 2.4 mL |

- Add 50 L of each Ractopamine Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

- Add 50 L of each sample in duplicate into different sample wells.

- Add 100 L of Antibody #1 and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 150 L of 1X Antibody #2. Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F) ( Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (F Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Ractopamine Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical ractopamine standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8, Moorestown,NJ08057

Tel.:856-727-0250         FAX.: 856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us