REAGEN™

Streptomycin ELISA Test Kit

RND99002

### Product Description

REAGEN™ Streptomycin ELISA Test Kit is a competitive enzyme immunoassay for the quantitative analysis of streptomycin and dihydrostreptomycin in feed, honey, kidney, liver, meat (beef, chicken and pork), milk, serum and urine. The unique features of the kit are:

- Rapid (10 - 30 minutes), and organic reagent-free extraction methods for various samples with high recovery (75- 115%).

- High sensitivity (0.5 ng/g or ppb) and low detection limit (5 ng/g or ppb for meat and milk).

- High reproducibility.

- A quick ELISA assay (less than 2 hours regardless of number of samples).

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Streptomycin ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C *. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| Streptomycin-coated Microtiter Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8C |
| Streptomycin Standards: Negative control (white cap tube) 0.5 ng/mL (yellow cap tube) 1.0 ng/mL (orange cap tube) 2.5 ng/mL (pink cap tube) 5.0 ng/mL (purple cap tube) 10 ng/mL (blue cap tube) 1000 ng/mL (spiking, red cap tube) | 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL 1.0 mL | 2-8C 2-8C |
| Streptomycin Antibody #1 | 15 mL | 2-8C |
| 100X HRP-Conjugated Antibody #2 | 250 L |  |
| Antibody #2 Diluent | 20 mL |  |
| 20X Wash Solution | 30 mL |  |
| Stop Buffer | 20 mL |  |
| TMB Substrate | 10 mL |  |
| Sample Balance Buffer 10X PBS 200X Sample Extraction Buffer | 0.5 mL 25 mL 25 mL |  |

*If you are not planning to use the kit for over 3 months, store Streptomycin Antibody #1 and 100X HRP-Conjugated Antibody #2 at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Feed | 5 |
| Honey | 10 |
| Meat/Liver/Kidney | 5 |
| Milk | 10 |
| Milk Powder | 25 |
| Serum | 5 |
| Urine | 5 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Streptomycin | 100 |
| Dihydrostreptomycin | 100 |
| Neomycin | <0.1 |
| Kanamycin | <0.1 |
| Gentamicin | <0.1 |
| Tobramycin | <0.1 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

### Warnings and Precautions

- The standards contain streptomycin. Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT-AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temperature (20 – 25C / 68 – 77F) or in a refrigerator before use.

- Preparation of 1X Sample Extraction

Mix 1 volume of 200X Sample Extraction Buffer with 199 volumes of distilled water.

- Preparation of 1X PBS

Mix 1 volume of the 10X PBS with 9 volumes of distilled water.

- Preparation of 1X PBS-Tween Buffer ( for milk sample)

Mix 1 volume of the 10X PBS with 9 volumes of distilled water, then add 0.1% Tween 20 (e.g. add 0.1 g of Tween 20 in 100 mL of 1-PBS).

### Feed

- To a reasonable amount of the feed sample, add 10 times the amount of 1X Sample Extraction Buffer (e.g. 1 g of sample + 10 mL of 1X Sample Extraction Buffer), homogenize the sample with a suitable mixer.

- Take out 1.2 mL of the homogenized sample and centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 0.5 mL of the supernatant to a new tube.

- Incubate the supernatant at 75°C for 5 minutes and vortex for 1 min at maximum speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 0.2 mL of the supernatant to a new tube, add 5 L of Sample Balance Buffer and mix well.

- Use 50 L per well for the assay.

Note: Dilution factor: 10. If needed, the sample obtained from step 6 can be further diluted with 1X PBS.

### Honey

- Add 1 g of the honey sample to 19 mL of 1X Sample Extraction Buffer, and mix well.

- Centrifuge 0.5 mL of the sample at 4,000 x g for 5 minutes.

- Transfer 0.2 mL of the supernatant to a new tube, add 5 L of Sample Balance Buffer and mix well.

- Use 50 L per well for the assay.

Note: Dilution factor: 20. If needed, the sample can be further diluted with 1X PBS.

### Meat/Liver/Kidney

- Remove fat from the sample. Homogenize the sample with a suitable mixer.

- Add 4.5 mL of 1X Sample Extraction Buffer to 0.5 g of the homogenized sample. Vortex for 3 minutes at maximum speed or rotorack for 20 minutes ( e.g. VWR Tube Rotator ).

- Take out 1.2 mL of the homogenized sample and centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 0.5 mL of the supernatant to a new tube (Avoid the fatty layer!).

- Incubate the supernatant at 75°C for 5 minutes and vortex for 1 minute at max speed.

- Centrifuge for 5 minutes at 4,000 x g at room temperature (20 – 25C / 68 – 77F).

- Transfer 0.2 mL of the supernatant to a new tube (Avoid the fatty layer!), add 5 L of Sample Balance Buffer and mix well.

- Use 50 L per well for the assay.

Note: Dilution factor: 10. If needed, the sample obtained from step 7 can be further diluted with 1X PBS.

### Milk

- For fat-free milk, dilute the milk sample with 1X PBS-Tween Buffer (1:19) (e. g. 20 L of milk + 380 L of 1X PBS-Tween Buffer). Take 50 L of the diluted sample per well for the assay.

- For the regular milk with fat, centrifuge the milk sample at 4,000 x g for 5 minutes, discard the upper fat layer. Dilute the sample with 1X PBS-Tween Buffer (1:19) (e. g. 20 L of milk + 380 L of 1X PBS-Tween Buffer). Take 50 L of the diluted sample per well for the assay.

Note: Dilution factor: 20. If needed, the sample can be further diluted with 1X PBS.

### Milk Powder

- Weigh out 2 g of the milk powder sample in a centrifugal vial, add 10 mL of distilled water and dissolve by shaking.

- Centrifuge at 4,000 x g for 5 minutes, discard the upper lipid layer.

- Dilute the sample with 1X PBS-Tween Buffer (1:9) (e. g. 20 L of the sample + 180 L of 1X PBS-Tween Buffer).

- Take 50 L of the diluted sample per well for the assay.

Note: Dilution factor: 50. If needed, the sample can be further diluted with 1X PBS.

### Serum

- Centrifuge 0.5 mL of the serum sample at 4,000 x g for 5 minutes.

- Take 20 L of the supernatant and add 180 L of 1X PBS, mix well.

- Use 50 L per well for the assay.

Note: Dilution factor: 10. If needed, the sample can be further diluted with 1X PBS.

### Urine

- Centrifuge 0.5 mL of the urine sample at 4,000 x g for 5 minutes.

- Take 20 L of the supernatant and add 180 L of 1X PBS, mix well.

- Use 50 L per well for the assay

Note: Dilution factor: 10. If needed, the sample can be further diluted with 1X PBS.

STREPTOMYCIN ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2 Diluent.

- Preparation of 1X Wash Solution

Mix 1 volume of the 20X Wash Solution with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| Streptomycin Antibody #1 | 100 L | 2.4 mL |
| 1X HRP-Conjugated Antibody #2 | 150 L | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

- Add 50 L of each Streptomycin Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

- Add 50 L of each sample in duplicate into different sample wells.

- Add 100 L of Antibody#1 and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F). 

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 150 L of 1X Antibody #2 to each well. Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F) (  Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (  Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Streptomycin Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

The following figure is a typical Streptomycin standard curve

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us