REAGEN™

Benzo[a]pyrene (BAP) ELISA Test Kit

RNC95006

### Product Description

REAGEN™ Benzo[a]pyrene (BAP) ELISA Test Kit provides a competitive enzyme immunoassay for the quantitative analysis of Benzo[a]pyrene in water, vegetable oil, fish, shrimp and meat . The unique features of the kit are:

- High recovery (>75%), Rapid and cost-effective extraction methods for various samples.

- High sensitivity (0.25ng/g or ppb)

- A quick ELISA assay (less than 2 hours regardless of number of samples).

- High reproducibility.

### Procedure Overview

The method is based on a competitive colorimetric ELISA assay. The drug of interest has been coated in the plate wells. During the analysis, sample is added along with the primary antibody specific for the target drug. If the target is present in the sample, it will compete for the antibody, thereby preventing the antibody from binding to the drug attached to the well. The secondary antibody, tagged with a peroxidase enzyme, targets the primary antibody that is complexed to the drug coated on the plate wells. The resulting color intensity, after addition of substrate, has an inverse relationship with the target concentration in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™ Benzo[a]pyrene (BAP) ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C*. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| BAP-coated Microtiter Plate | 1 x 96-well plate (8 wells x 12 strips) | 2-8C |
| 1000PPb BAP Standards stocks | 0.8 mL | 2-8C |
| BAP Antibody #1 | 12 mL | 2-8C * |
| 100X HRP-Conjugated Antibody #2 | 250 L | 2-8C * |
| Antibody #2 Diluent | 20 mL | 2-8C |
| 5X Sample Suspension Buffer | 25 mL | 2-8C |
| 20X Wash Solution ** | 28 mL | 2-8C |
| Stop Buffer | 20 mL | 2-8C |
| TMB Substrate | 12 mL | 2-8C |
| BAP clean up mix | 10g | 2-8C |

* 	If you are not planning to use the kit for over 3 months, store BAP Antibody #1 and 100X HRP-Conjugated Antibody #2 at -20C or in a freezer.

### Sensitivity

| Sample Type | Detection Limit (ng/g or ppb) |
| --- | --- |
| Shrimp/Fish/Meat | 2.5 |
| Vegetable Oil | 12.5 |
| Water | 0.3 |

### Specificity

| Analytes | Cross-Reactivity (%) |
| --- | --- |
| Benzo[a]pyrene | 100 |
| Hexachlorobenzene | <0.1 |
| Hexachlorocyclopentadine | < 0.1 |
| Methoxychlor | < 0.1 |

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Tissue Mixer (e.g. Omni TissueMaster Homogenizer)

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Evaporator (Optional)

- MgSO4.

### Warnings and Precautions

- The standards contain Benzo[a]pyrene . Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT- AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

# SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

Preparation of 1X Sample Extraction Buffer:

Mix 1 volume of 5X Sample Extraction Buffer with 4 volumes of distilled water.

### Fish/Shrimp/Meat

- Take 2g of the homogenized sample,add 6.0 mL of acetonitrile, and add 1.0 g of anhydrous MgSO4 ,vortexing for 3 mins.

- Centrifuge at 4,000 x g for 10 minutes at room temperature (20 – 25C / 68 – 77F).

- Take 0.5 mL of supernatant to a new 2-mL tube containing 125 mg of BaP Clean Up Mix

- Vortex vigorously for 1 minute and leave at room temperature for 5 minutes. Centrifuge the sample for 10 minutes at 4,000 X g at room temperature (20-25 oC/68-77oF).

- Take 200 μL of supernatant to a 2-mL glass tube. Use an evaporator to dry the sample at 50C under reduced pressure in the dark. Alternatively, the sample can be dried by blowing nitrogen gas in a 50C water bath or at room temperature for a longer time. (Note: Avoid over drying the sample. If the sample is over dried and "hard crack" appears in the tube, the sample can not be resuspended well and the recovery rate could be reduced).

- Add 500 μL of 1X BaP Sample Suspension Solution to the dried tube and vortex vigorously for 2 minutes.  Alternatively, to fully dissolve the dried residue, vortex vigorously for 1 minute following by sonicating the sample for 5 minutes at maximum force in a water bath sonicator.

- Use 50 L of the lower aqueous layer per well for the assay.

Note: Dilution factor: 10. (fit for 5-20 ppb spiking). To detect the higher concentration BaP in the sample, the dilution factor can be increased by adding more 1X BaP Sample Suspension Solution in Step 8.  For example, if add 800 μL 1X BaP Sample Suspension Solution, the dilution factor is 20 (fit for 10-40 ppb spiking).

### Vegetable Oil

- Add 1 mL 90% N,N-dimethylformamide (DMF) (DMF/H2O, 9:1 V/V) and 3 mL n-hexane into 1 mL

vegetable oil. Vortex vigorously for 10 minutes in a multi-tube vortexer or shaker or vortex for 3

minutes at maximum speed manually.

- Centrifuge the sample for 10 minutes at 4,000 X g at room temperature (20-25°C/68-77 °F).

- Remove the top hexane layer completely.

- Take 20 μL bottom DMF solution, dilute with 980 μL of 1X BaP Sample Suspension Solution

Note: Dilution factor: 50  If the BaP concentration is expected to be lower (5-50 ppb), Step 4 above can be changed as following: “Take 50 μL bottom DMF solution, dilute with 950 μL of 1X BaP Sample Suspension Solution.” The dilution factor will be 20.

### Sea water

- Take 1.5 mL of water sample in a centrifuge tube and adjust the pH to 6.5-7.5 with 0.1 M HCl or 0.1 M NaOH.

- Centrifuge the water sample for 5 minutes at 4,000 X g at room temperature (20-25 oC/68-77oF).

- Take out 0.8 mL of the supernatant to a new tube, add 0.2 mL of 5X BaP Sample Suspension Solution. Vortex 30 seconds.

- Use 50 μL of the solution per well in the ELISA test.

Note: Dilution factor: 1.25.

### Instant  Noodles

- Take 2g of the homogenized sample,add 8.0 mL of acetonitrile,vortexing for 1 mins and ultrasonic extract for 5 mins.

- Centrifuge at 4,000 x g for 10 minutes at room temperature (20 – 25C / 68 – 77F).

- Take 0.5 mL of supernatant to a new 2-mL tube containing 125 mg of BaP Clean Up Mix

- Vortex vigorously for 1 minute and leave at room temperature for 5 minutes. Centrifuge the sample for 10 minutes at 4,000 X g at room temperature (20-25 oC/68-77oF).

- Take 200 μL of supernatant to a new glass tube. Use an evaporator to dry the sample at 50C under reduced pressure in the dark. Alternatively, the sample can be dried by blowing nitrogen gas in a 50C water bath or at room temperature for a longer time. (Note: Avoid over drying the sample. If the sample is over dried and "hard crack" appears in the tube, the sample can not be resuspended well and the recovery rate could be reduced).

- Add 500 μL of 1X BaP Sample Suspension Solution to the dried tube and vortex vigorously for 2 minutes.

- Use 50 L of the lower aqueous layer per well for the assay.

Note: Dilution factor：10

Benzo[a]pyrene ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 100X HRP-Conjugated Antibody #2 with 99 volumes of Antibody #2 Diluent.

- Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Buffer concentrate with 19 volumes of distilled water.

- Preparation of Benzo[a]pyrene Standard

Mix the following solutions by serial dilution starting from the highest concentration to the lowest in 2-mL microcentrifuge tubes.

Make standards fresh each time performing the assay.

| Work Standards | Volume of BaP Source | Volume of 1x BaP Suspension Solution |
| --- | --- | --- |
| 25 ppb | 25 µL (of 1000 ppb stock) | 975 µL |
| 2.5 ppb | 100 µL (of 25 ppb dilution) | 900 µL |
| 1.5 ppb | 600 µL (of 2.5 ppb dilution) | 400 µL |
| 1.0 ppb | 500 µL (of 1.5 ppb dilution) | 250 µL |
| 0.5 ppb | 500 µL (of 1.0 ppb dilution) | 500 µL |
| 0.25 ppb | 500 µL (of 0.50 ppb dilution) | 500 µL |
| 0 ppb | 0 µL | 500 µL |

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction | 24 Reactions |
| --- | --- | --- |
| BAP Antibody #1 | 100 L | 2.4 mL |
| 1X HRP-Conjugated Antibody #2 | 150 L | 3.6 mL |
| 1X Wash Solution | 2.0 mL | 48 mL |
| Stop Buffer | 100 L | 2.4 mL |
| TMB Substrate | 100 L | 2.4 mL |

- Add 50 L of each BAP Standards in duplicate into different wells ( Add standards to plate only in the order from low concentration to high concentration).

- Add 50 L of each sample in duplicate into different sample wells.

- Add 100 L of BAP Antibody #1 and mix well by gently rocking the plate manually for 1 minute.

- Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 150 L of 1X Antibody #2 solution. Incubate the plate for 30 minutes at room temperature (20 – 25C / 68 – 77F) (Avoid direct sunlight and cold bench tops during the incubation. Covering the microtiter plate while incubating is recommended).

- Wash the plate 3 times with 250 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

- Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix the solution by gently rocking the plate manually for 1 minute while incubating (Do not put any substrate back to the original container to avoid any potential contamination. Any substrate solution exhibiting coloration is indicative of deterioration and should be discarded. Covering the microtiter plate while incubating is recommended).

- After incubating for 15 minutes at room temperature (20 – 25C / 68 – 77F), add 100 L of Stop Buffer to stop the enzyme reaction.

- Read the plate as soon as possible following the addition of Stop Buffer on a plate reader with 450 nm wavelength ( Before reading, use a lint-free wipe on the bottom of the plate to ensure no moisture or fingerprints interfere with the readings).

### Benzo[a]pyrene (BAP) Concentration Calculations

A standard curve can be constructed by plotting the mean relative absorbance (%) obtained from each reference standard against its concentration in ng/mL on a logarithmic curve.

Relative absorbance (%) =

Use the mean relative absorbance values for each sample to determine the corresponding concentration of the tested drug in ng/mL from the standard curve.

The following figure is a typical Benzo[a]pyrene (BAP) standard curve.

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250         FAX.: 856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.: www.reagen.us