REAGEN™

E. coli O157:H7 ELISA Test Kit

RNB93002

### Product Description

REAGEN™E. coli O157:H7 ELISA Test Kit is designed for routine screening for E. coli O157H7 from cultured food and feed samples. E. coli serotype O157:H7 is a gram-negative rod-shaped bacterium.The letter "O" in the name refers to the somatic antigen number, whereas the "H" refers to the flagella antigen. This is one of hundreds of serotypes of the bacterium E. coli. While most strains are harmless and normally found in the intestines of mammals, this O157:H7 strain may produce Shiga-like toxins,cause severe illness, and is a member of a class of pathogenic E. coli known as enterohemorrhagic E.coli or EHEC. People infected with enterohemorrhagic E. coli often develop hemorrhagic colitis and/or hemolytic-uremic syndrome. The sources of contamination include raw ground beef, unpasteurized milk,untreated water, juice, raw sprouts, lettuce, and salami.

The detection of E. coli O157:H7 in food samples is often laborious and time-consuming. Several days are required to culture the bacteria from a sample followed by a final isolation and identification by biochemical and serological techniques. The REAGEN™ E. coli O157:H7 ELISA Test Kit provides a simple, rapid, sensitive and cost-effective enzyme immunoassay (ELISA) screening method, which enables government agencies, seafood processors, as well as quality assurance organizations, to detect positive and negative E. coli O157:H7 samples to as low as 105 cells/mL level.

### Procedure Overview

The method is based on a colorimetric ELISA assay. The monoclonal antibody against E. coli O157:H7 has been coated in the plate wells. During the analysis, the sample is added. If E. coli O157:H7 is present in the sample, it will bind with the antibody. The secondary antibody, tagged with a peroxidase enzyme, targets the E. coli O157:H7 that is complexed to the primary antibody coated on the plate wells.The resulting color intensity, after addition of substrate, is related to the amount of E. coli O157:H7 in the sample.

### Kit Contents, Storage and Shelf Life

REAGEN™E.coli.0157:H7 ELISA Test Kit has the capacity for 96 determinations or testing of 42 samples in duplicate (assuming 12 wells for standards). Return any unused microwells to the foil bag and reseal them with the desiccant provided in the original package. Store the kit at 2-8C*. The shelf life is 12 months when the kit is properly stored.

| Kit Contents | Amount | Storage |
| --- | --- | --- |
| E. coli O157:H7 antibody-coated Plate | 96well plate(8wells x12 strips) | 2-8C |
| 1X E. coli O157:H7 Negative Control | 1 mL | -20C |
| 1X E. coli O157:H7 Positive Control | 1 mL | -20C * |
| 50X HRP-Conjugated Antibody #2 | 50 Lx3 | -20C * |
| Antibody #2 Diluent ** 20X Wash Solution ** Stop Buffer ** TMB Substrate ** | 20 mL 30 mL 20 mL 10 mL | 2-8C |

### Sensitivity

The kit can detect 105cells/mL.

### Specificity

The kit can broadly detect E. coli O157:H7 antigen.

### Required Materials Not Provided With the Kit

- Microtiter plate reader (450 nm)

- Incubator

- Vortex mixer (e.g. Gneie Vortex mixer from VWR)

- 10, 20, 100 and 1000 L pipettes

- Multi-channel pipette: 50-300 L (Optional)

- Modified TSB Broth (mTSB): Consisting of Trypticase soy broth (TSB) supplemented with Casamino Acids (10 g/liter) and K2HPO4 (1.5 g/liter)

### Warnings and Precautions

- The standards contain E.coli.0157:H7 . Handle with particular care.

- Do not use the kit past the expiration date.

- Do not intermix reagents from different kits or lots except for components with the same part No’s within their expiration dates. ANTIBODIES AND PLATES ARE KIT- AND LOT-SPECIFIC. Make sure that the antibody #2 and diluent are mixed in correct volumes.

- Try to maintain a laboratory temperature of 20°–25°C (68°–77°F). Avoid running assays under or near air vents, as this may cause excessive cooling, heating and/or evaporation. Also, do not run assays in direct sunlight, as this may cause excessive heat and evaporation. Cold bench tops should be avoided by placing several layers of paper towel or some other insulation material under the assay plates during incubation.

- Make sure you are using only distilled or deionized water since water quality is very important.

- When pipetting samples or reagents into an empty microtiter plate, place the pipette tips in the lower corner of the well, making contact with the plastic.

- Incubations of assay plates should be timed as precisely as possible. Be consistent when adding standards to the assay plate. Add your standards first and then your samples.

- Add standards to plate only in the order from low concentration to high concentration as this will minimize the risk of compromising the standard curve.

- Always refrigerate plates in sealed bags with a desiccant to maintain stability. Prevent condensation from forming on plates by allowing them equilibrate to room temperature (20 – 25C / 68 – 77F) while in the packaging.

SAMPLE PREPARATION

Be sure samples are properly stored. In general, samples should be refrigerated at 2-4°C for no more than 1-2 days. Freeze samples to a minimum of -20°C if they need to be stored for a longer period. Frozen samples can be thawed at room temps (20 – 25C / 68 – 77F) or in a refrigerator before use.

### Meat

25 g of meat sample is stomached for 2 minutes and cultured overnight at 37°C with agitation at 120 rpm in 225 mL of mTSB medium. Take out 1 mL culture, autoclave at 121°C or boil for 20 minutes. Cool the sample to room temperature. Mix and perform the MaxSignal® E. coli O157:H7 ELISA Test.

E.coli.0157:H7 ELISA TEST KIT PROTOCOL

### Reagent Preparation

IMPORTANT: All reagents should be brought up to room temperature before use (1 – 2 hours at 20 – 25C / 68 – 77F); Make sure you read “Warnings and Precautions” section on page 3. Solutions should be prepared just prior to ELISA test.  All reagents should be mixed by gently inverting or swirling prior to use. Prepare volumes that are needed for the number of wells being run. Do not return the reagents to the original stock tubes/bottles. Using disposable reservoirs when handling reagents can minimize the risk of contamination and is recommended.

- Preparation of 1X HRP-Conjugated Antibody #2

Mix 1 volume of 50X HRP-Conjugated Antibody #2 with 49 volumes of Antibody #2 Diluent.

- Preparation of 1X Wash Solution

Mix 1 volume of 20X Wash Buffer concentrate with 19 volumes of distilled water.

### ELISA Testing Protocol

Label the individual strips that will be used and aliquot reagents as the following example:

| Component | Volume per Reaction |
| --- | --- |
| 1X Negative Control | 50 L |
| 1X Positive Control | 50 L |
| 1X HRP-Conjugated Antibody #2 | 50 L |
| 1X Wash Solution | 2.0 mL |
| Stop Buffer | 100 L |
| TMB Substrate | 100 L |

1.  Take out the needed plate strips, add 50 L of Negative and Positive Control in duplicate into different wells.

2.  Add 50 L of each sample in duplicate into different sample wells.

3.  Incubate the plate for 30 minutes at room temperature (20 – 25 C / 68 – 77 F).

4.  Wash the plate 3 times with 300 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels ( Perform the next step immediately after plate washings. Do not allow the plate to air dry between working steps).

5.  Add 50 L of 1X HRP-Conjugated Antibody #2 to each well. Incubate the plate for 30 minutes at room temperature (20 – 25 C / 68 – 77 F)

6.  Wash the plate 3 times with 300 L of 1X Wash Solution. After the last wash, invert the plate and gently tap the plate dry on paper towels .

7.  Add 100 L of TMB substrate. Time the reaction immediately after adding the substrate. Mix thesolution by gently rocking the plate manually for 1 minute while incubating .

8.  After incubating for 15 minutes at room temperature (20 – 25 C / 68 – 77 F), add 100 L of StopBuffer to stop the enzyme reaction.

9.  Read the plate as soon as possible following the addition of Stop Buffer on a plate reader wit450nm wavelength .

Interpretation of the Results

REAGEN LLC

Add.: 102CommerceDr.Unit8,Moorestown,NJ08057

Tel.:856-727-0250        FAX.:856-727-0251

E-mail:reagenkits@gmail.com  reagenllc@gmail.com

Web.:www.reagen.us

### Visual Interpretation of the Results

Visual assessment must be done prior to addition of stopping solution.

Positive Control:     Blue

Negative Control:     Clear or very pale blue

A sample is considered positive when the following criteria are met:

The negative control is clear or very pale blue in color and the positive sample is significantly deeper in color than this.

The positive sample should always be significantly deeper in color than the negative control, however it need not be as blue as the positive control.

A sample is considered negative when the following criteria are met:

The negative control is clear or very pale blue in color.

The sample is less than or equal to the color of the negative control.

Interpret results and record as positive or negative on the record sheet provided. Return reagents to 2-8°C. Negative/Positive samples can only be interpreted if the expected results are produced with the controls.

### OD Method

Measure the plate absorbance using a microplate reader at A450nm and record absorbance values on the record sheet provided.

Negative Control:    Readings at A450nm should give an OD of less than 0.4.

If not this indicates insufficient washing of the microplate.

Positive Control:    Readings at A450nm should give an OD greater than 1.0.

If controls are not within the limits noted above, the test must be repeated.

Negative samples:    Samples that give A450nm <0.6

Positive samples:    Samples that give A450nm >0.6

Negative/positive samples can only be interpreted if the expected results are produced with the controls.Samples positive by the MaxSignal  E. coli O157:H7 ELISA Test.should be confirmed from unheated ELISA-5-bromo-4-chloro-3-indolyl-b-D-glucuronide and then eosin methylene blue (EMB) agar. Putative E. coli O157:H7 isolates can be tested for agglutination with antiserum to the O157 antigen (Difco).E.coli.0157:H7 (AOZ) Concentration Calculations